Developing a Polarimeter to determine the optical
constants of materials in the XUV

Raymond Rios
Research Advisor: Dr. Steve Turley

Abstract: This research paper discusses the progress made in developing a XUV polarimeter, a method for determining the optical constants of in the range of light with wavelengths measuring between 111m to 50ilm. Current methods of determining these constants with ellipsometry have been limited in their ability due to their high absorption coefficients in this wavelength region. The Polarimeter shows intensity variations using both Channeltron and CCD camera detectors due to the rotation of two mirrors relative to each other. It is hoped that these variations are consistent with predicted changes due to polarization which would lead, through mathematical modeling, to the optical constants of the mirrors. Also included in this paper are operating and alignment procedures for the polarimeter.

Table of Contents

1.	Introduction

i.	Background

ii.	Project objectives

iii.	Project accomplishments

2. 	Construction and Testing of Polarimeter

i.	Channeltron verses CCD camera

ii.	Optimizing Channeltron

iii.	Construction of adjustable aperture

iv.	Calculating the expected number of counts for the channeltron

v.	Procedural techniques and complications with Channeltron

vi.	Procedural techniques and complications with CCD camera

3. 	Data Results

i.	Channeltron detector

ii.	CCD camera

4.	Conclusions

5.	Appendix A: Standard Operating Procedures

i.	Taking monochrometor to vacuum

ii.	Turning on Source

iii. 	Taking O-chamber and Polarimeter to vacuum

iv. 	Making measurements with a Channeltron detector

v.	Making measurements with a CCD camera

vi. 	Shutting off Source and bring up monochrometor to roughing

vii.	Bringing O-chamber and Polarimeter to ATM

6.	Appendix B: Alignment Procedures

i.	Alignment procedures for the Polarimeter pinhole

ii.	Alignment procedures for the Channeltron detector

7.	Appendix C: Project Accomplishments

i.	Work Done on Project ii. Data Analysis Results

ii.	Miscellaneous correlating research work

Table of Images

Figure 1: Polarimeter with channeltron detector attached to O-chamber

Figure 2: Simplified Polarimeter design

Figure 3: Possible view of light beam movement after passing through the polarimeter Figure 4: Adjustable Pinhole design

Figure 5: Polarimeter measurement setup using the channeltron detector

Figure 6: Entrance to the Polarimeter taken from inside the O-chamber

Figure 7: Polarimeter measurement setup using the CCD camera

Figure 8: Polarimeter measurements using the channeltron detector

Figure 9: Polarimeter measurements using the CCD detector with background

interference.

Figure 10: Background with Background light

Figure 11: Background with background light diminished

Figure 12: Alignment technique for the pinhole to the polarimeter.

Figure 1: Polarimeter with channeltron detector attached to 0-chamber

REU Summer Research Goal:

Development of a polarimeter to determine the optical properties of materials in the extreme ultraviolet.

Background:

Over the past two years, the Brigham Young University physics department has been examining the development of a polarimeter used to find the reflectance of materials in the XT-TV range of light (10 to 40 eV in the electromagnetic spectrum). . Because sensitivity and difficulty of measuring material properties in this range with previous methods, the optical properties of materials in the XUV

have been are relatively uncertain. To find the optical properties of materials, it is important to know some of the properties of the XUV source that is used in the measurement. Using monochromatic and polarized light, we restrict the amount of variables that can influence the reflection of the light. BYU has a monochrometor that provides a relatively steady monochromatic light source.

Unfortunately the monochrometor does not polarize the light. In order to eliminate the need for polarized light, we have built a polarimeter that does not require the light to be polarized to find the optical properties of the material. The polarimeter uses two mirrors of the same material to do this as shown in figure 2. Unpolarized light enters the polarimeter and reflects off the first mirror at 45degree angle. This causes the light to be polarized in one direction. The light then strikes the second mirror with 45 degree with the incident beam. We are then able

to rotate the second mirror along the direction the light strikes the second mirror.

By finding the relative intensities at different angles, we can then take that

information and model it on a computer to determine the possible optical

constants that would give us those results.

Figure 2: Simplified Polarimeter design

Research this summer:

This summer I will test the polarimeter with materials that have accepted values in the XUV and then compare them to determine the effectiveness of the polarimeter. If time permits, I will test other materials that will be useful for development in thin multilayer mirrors that can be used for space, lithography and XUV laser applications in the future.

Channeltron Versus CCD:

The Polarimeter was designed to be used by either a CCD camera or Channeltron detector. Because both the channeltron and the CCD camera have they're own distinct advantages in taking measurements, part of this research was to determine which detector was best suited for the polarimeter.

The number of photons recorded by the channeltron is summed up over the area it's detection area. This light sensitive area on the channeltron is called the CEM. The CCD camera uses a matrix of extremely small optically sensitive

semiconductors. Unlike the CCD camera, which can provide thousands of data

points for one measurement, the channeltron provides only one. When compared

to the CCD camera, the channeltron detector has some benefits as well as some

drawbacks.

While the CCD camera has more data points, the channeltron only has one which decreases the processing time required by the computer and hence the overall measurement duration. This is important because we find that the light source to be unstable and has a tendency to fluctuate in intensity. Since the overall

measurement will be shorter with the channeltron than the CCD camera, the beam intensity is less likely to wander as far compared to the amount of time it takes to make measurements with the CCD camera. Also, since there is only one data point, the data can be taken directly as opposed to determining the relative number of photons that hit the CCD camera by having to process each image in Matlab or some other program.

Unfortunately, the channeltron does not provide any other information about the beam itself. The CCD camera, while it does take longer to record and process the results from the measurement, gives the user much more information about the beam. For example: beam image, position and uniformity.

Figure 3: Possible view of light beam movement after passing through the polarimeter

Optimizing Channeltron:

We determined that the channeltron is best suited for measurements done by the polarimeter so long as two requirements are met: One, the beam position

remained either stationary or uniform and two, the surface area of the CEM was large enough to collect enough counts to analyze. In using the channeltron we have two options:

1. Completely cover the entrance hole with the beam

2. Limit the beam size so that all of it enters the entrance hole

Only having part of the beam enter the hole will not do because by rotating the

beam even with small fluctuations in the beams position can alter greatly the

relative number of counts and through off the measurements.

By covering the entrance hole completely we get the most counts possible. Plus the beam has more flexibility if it moves while rotating the second mirror. The one concern is that if beam is not uniform, then as the mirror is rotated the intensity will vary do to both polarization and changes in the beam intensity.

When the beam is limited to within the diameter of the detector hole then variations of intensity with in the beam do not matter since the entire beam enters the entrance hole. Limiting the beam has its own problems though. The number of counts drops, which, when dealing with reflection off of two mirrors in the XUV range, is a major concern. The other problem is that the beam may wander and not hit the detector completely. To keep the beam from wandering, a movable pinhole is positioned so that the beam strikes the center of rotation of both mirrors and thus keeps the beam's position stationary. The tolerance in the beam's movement is limited by the size of the beam. The larger the beam means more counts but less flexibility and more likelihood that the beam moves slightly out of the entrance hole.

If the beam intensity is uniform over the entire width, then the best solution is to completely cover the entrance slit. One concern may be that by making the beam too large there might be internal reflection inside the polarimeter that still makes it into the detector. Such effects on the measurements should be minimal since the intensity of the beam dies off dramatically with each reflection. If there remains any concern, the beam should remain large enough to cover the entrance hole with flexibility in movement yet small enough to keep the beam away from the walls of the polarimeter and the size of the beam.

Construction of Adjustable Aperture:

During the initial course of this research, the uniformity of the beam was unknown. There was a possibility that the concentration of the beam laid in the center and the diminished radially outward. The alignment of the detector was crucial to make certain that it captured only the light that reflected at the center of rotation. Set screws were used to allow the pinhole to be adjustable. The diagram in figure 4 shows the design used for the construction

of the pinhole.

Procedures for aligning the polarimeter pinhole and detector see Appendix B.

Calculating the expected amount of counts through polarimeter:

It is important to know what the expected amount of counts should be when determining whether or not the polarimeter is working properly. This can be accomplished by taking the number of counts made by the channeltron detector within the O-chamber a small distance before the beam strikes the polarimeter aperture. By comparing the relative sizes of the entrances to the polarimeter and detector and multiplying them by the percentages of reflection at 45-degrees from two mirrors, the number of counts will be scaled to the expected output value of the detector. We can then compare this result to actual measurements to see if there are any major discrepancies. The equation for this is given by,

Where,

Np = Number of counts predicted after passing through the polarimeter

Nd = Number of counts measured with detector before the polarimeter

Ap = Cross sectional area of the polarimeter aperture

Ad = Cross sectional area of the detector aperture

rl = Percentage of reflection off of the first mirror

r2 = Percentage of reflection off of the second mirror.

Using this method, 145,000 counts in 0.5 second intervals before the polarimeter should give about 4,500 counts for the same amount of time after passing through the polarimeter.

Procedural techniques and complications with Channeltron:

All measurements using the Channeltron and CCD camera were done with 584nm wavelength light. Figure 5 shows the lab set up using the channeltron detector.

1. Gas tank

2. XUV source

3. Monochrometor with diffraction grating

4. O-chamber

5. Polarimeter

6. Channeltron detector inside casing

7. Output coaxial cable adapter

8. Computer interface box

9. Detector power source

10. Detector power switch with reset button

11. Interlock box. Contains detector manual override switch

For procedures on making measurements with the Channeltron detector see Appendix A

Procedural techniques and complications with CCD camera:

1. Gas tank

2. XUV source

3. Monochrometor with diffraction grating

4. 0-chamber

5. Polarimeter

6. CCD camera

7. Liquid coolant controller

8. Liquid coolant pump

9. Computer

10. CCD camera cooling controller

For procedures on making measurements with the Channeltron detector see Appendix A

Channeltron detector:

The number of counts recorded by the channeltron detector was sadly smaller than expected. Figure 7 shows the number of counts obtained with the second mirror set at different angles (delta).

According to estimates made by the calculations shown above the number of counts should be five times larger at zero. While the above curve demonstrates the expected motion of lowering intensity due to polarization effects, the

difference between the initial and final angle is much larger than what IMD predicts. Furthermore when the program for determining the optical constants was applied using this data, the resulting answer was not anyway near accepted values. We must conclude that either the program is faulty or that there is something wrong with the data obtained with the channeltron.

Pinhole alignment with source:

Figure 11: Entrance to the Polarimeter taken from inside the O-chamber

The most likely cause for the low number of counts, besides the detector alignment, is due to the alignment of the pinhole. The pinhole to the polarimeter is aligned so that the beam hits the center of rotation of the second mirror. Due to construction of the polarimeter, the pinhole may have to be set some position off center as in this case shown in figure 11. This unfortunately has negative consequences. There are two large possibilities why the beam may not be as intense as predicted. One, because the pinhole is off center with the beam, the intensity may have dramatically decreased. The second, any light that enters the pinhole may have to do so at an angle. The channeltron detector is aligned so that light coming in normal to the pinhole will enter the center of the detector. If the light enters into the polarimeter at an angle other than normal, the beam would hit some other spot.

It is possible that because the beam enters at an angle the beam no longer hits the center of rotation and therefore moves when the second mirror is rotated. The beam may then not completely enter the detector and cause the intensity to decrease more than it normally would if stayed in the same position. To see if the problem is due to the channeltron. this same measurement was done using a CCD camera.

CCD camera:

While taking pictures of the XUV beam with the CCD camera, strange images appeared that gave concern as to what was happening within the polarimeter.

Figure 8: Polarimeter measurements using the CCD detector with background interference.

Figure 8 shows images taken at adjusted angles using a CCD camera with exposure of two minutes. The background mask is taken with the polarimeter set at zero. Assuming that natural light has no effect on the camera, it appears that

there may be some internal reflection. Upon further analysis, though, it was discovered that there was a large amount of noise due to the natural light that entered into the O-chamber and then through the polarimeter. This noise was easily brought down when the windows to the O-chamber were covered with felt cloth. The amount of noise was brought down to the same magnitude as the XUV image.

Figure 9: Background image with Background light	Figure 10: Background image with background light diminished

The image in figure 9 shows the light at zero degrees without covering the window portals on the O-chamber. The image of the XUV beam is not resolvable unless the light a background image is taken and then subtracted from this image. The first image in figure 8 has a close up image of what the beam would then look like.

The image in figure 10 shows the background light with the XUV beam clearly visible once the image is scaled down in intensity. By accounting for this by subtracting the background noise a clear image results. This however must be

done for each angle and results in doubling the amount of time needed to make a complete set of measurements. Measurements such as this can be done with both the portals covered or uncovered. They will both show images such as found in the first image in figure 8 at all angles; however, by covering the portals, the image resolution is sharper at higher angles.

The images produced by the CCD camera show that the beam intensity is uniform over the entire area. It also shows that the diameter of the beam is about 2.6mm which is small enough to enter the 3mm diameter channeltron detector hole.

Conclusions:

Appendix A

Standard Operating Procedures:

Taking monochrometor to vacuum:

Make sure Butterfly valve between monochrometor and O-chamber is closed and the source valve is open

Check to see if roughing pump is running

Check to see if cooling water is flowing

Turn turbo pump on

Turn thermal couple gauge on

Wait for turbo pump to reach normal operating speed

Turn on cold cathode gauge

Turning on Source:

Open gas line valve thermal couple gauge should read between 500-900 mmhg

Wait 15-20 minutes or longer before turning on source (this appears to keep the source more stable once the power is on)

Check to see if cooling water is flowing

Check to see if monochrometor is under vacuum Check to see that source valve is open

Turn on voltmeter

Flip circuit breakers on from left to right

Press Green button power on

Turn voltage knob to 40, voltage should be around 700 Volts on gauge

Current should jump up. If not close source valve and then open once it does

Current should read around 0.50 A

Taking O-chamber and Polarimeter to vacuum:

Butterfly valve closed

Argon line valve closed

Gate valve closed

Vacuum cap on O-chamber in place

Turn on roughing pump

Open gate valve

Turn on turbo pump

Wait until turbo pump reaches normal operating speed

Turn on cold cathode gauge

Making measurements with a Channeltron detector:

Connect coaxial cable from computer to the output box next to the Polarimeter

Check to see that there is no obstruction within the O-chamber. (Stage and 0chamber detector are out of the way of the beam's path)

Open up the channeltron program Enter wavelength setting shown on monochrometor.

Set time duration to 1 or 2 seconds

Set data collection to continuous

Make sure that the O-chamber is below 5x10^-5 Torr before turning on detector

Turn on power box to detector

Flip override switch labeled "detector"

Flip Detector switch next to the cold cathode gauge on

Push "Collect data" on channeltron program

Push reset button on

Open Butterfly valve

Wait for the output to steady then record in note book next to the angle and also the range at which the counts are reading

Rotate Polarimeter and repeat set 13

Once finished close Butterfly valve

Turn off all three detector power switches

Making measurements with a CCD camera:

With the CCD camera connected to the back of the Polarimeter and under vacuum, turn on liquid coolant. It has two switches on to the pump and on to the control box

Push index until you see the LED label read CL

Set CL to 20C by pushing the arrow buttons. (This is what room temperature should be around. The control box to this cooling system is not calibrated and resets to some random number each time it is turned on. It needs to be at a higher temperature than your goal so that it will cool the liquid coolant.)

Push index until you see A1

Set Al to some number above 20C. (This is Alarm one and the LED indicator should light up if the temperature is below this setting however it doesn't work like that and lights up if the temperature is above this setting. I don't know if it effects the cooling but I set it above my temperature so as not to cause any concern. Precise cooling temperature is not necessary since this is just the supporting cooling system)

Make sure goal is at -20C

Turn on controller located beneath the computer monitor Turn cooling knob to -50C

Turn on cooling switch on controller

Once the green light is on then the CCD camera is at that temperature

Seal off all window portals on O-chamber to light

Open Winview on windows 98 computer

Go to "Environment", "Setup" In the label box "name of sample +time interval + angle + `back"'

Set time interval to desired time (between 2 and 10 minutes)

Under the "Data Corrections" tab deselect background

Make sure butterfly valve is closed

Push "Acquire"

Once finished, go to "Environment", "Setup" In the label box "name of

sample +time interval + angle"

Under the "Data Corrections" tab select background and name "name of

sample +time interval + angle + `back"' (same name as was in step 12)

Open butterfly valve

Push "Acquire"

After it's done, close butterfly valve.

Repeat steps 12 to 22 for different angles keeping in mind the need to make

measurements for all the angles in one session and as quickly as possible

to keep conditions with the source relatively consistent.

Once finished, turn off all cooling switches

Turn off controller

Make sure butterfly valve is closed.

Give the CCD camera time to come back up to temperature before bringing 0

chamber to atmosphere.

Shutting off Source and bring up monochrometor to roughing:

Butterfly valve closed

Source valve open

Turn off cold cathode gauge

Turn voltage knob on power source to zero

Push red reset button

Flip circuit breaker switches off from right to left

Close off gas valves

Turn off Thermal couple gauge

Turn off Voltmeter

Turn off angle alignment motor

Push "Stop" on turbo pump

Leave roughing pump on

Bringing 0-chamber and Polarimeter to ATM:

Close gate valve.

Butterfly valve closed

Turn off Turbo pump on interlock box

Turn off roughing pump on interlock box

Turn on Argon switch on interlock box

Open up gas valves to Argon tank

Open gas valve above o-chamber

Wait until the vacuum cap can be removed easily, but leave on.

Close gas valves to Argon tank

Close gas valve above O-chamber

Turn off Argon switch on interlock box

Appendix B

Alignment procedures for the Polarimeter pinhole: (refer to image 12)

Without any detectors mounted on the polarimeter, place grid paper on the exit hole of the polarimeter. (Using tape is fine, it just needs to be cleaned off before placing it in a vacuum)

Bolt a transparent plate over the entrance of the polarimeter

Secure the polarimeter to the table using clamps

Raise M2 to the same height as the entrance hole and M1 to the same height as the laser.

Rotate M1 so that the laser hits M2. Rotate M2 to hit the polarimeter. (As of yet it is not so important that the beam hits the hole. It just needs to be close)

While looking at the polarimeter using M3 for convenience, adjust the beam with M1 so that it hits the center of the entrance hole.

Adjust M2 so that the reflected beam goes back the same way it came and hits the laser where the beam initial comes out. (Do not worry that the beam might no longer hit the entrance hole of the polarimeter. One suggestion: Try to see where the reflected beam hits M2 and adjust the mirror so that both the initial and reflected beams hit the same spot on M2. Then look at the laser and see how close it comes to the laser source.)

Repeat steps 5 through 7 until it is no longer necessary.

Rotate the polarimeter. If the entrance hole is in the proper place the beam will not move. If the beam spot does move, it will move a quarter of a circle. Adjust the set screws for the pinhole, to position it where the center of rotation appears to be.) Repeat steps 5-8.

Figure 12: Alignment technique for the pinhole to the polarimeter.

Alignment procedures for the Channeltron detector:

Place the detector mount without the detector onto the polarimeter

Take a piece of clear tape and place it over the exit hole

Rotate the polarimeter and check to see if the spot moves.

If the spot does move, make sure that the laser is aligned properly by checking

to see if the beam returns back to the laser. (This assumes that you still

have the transparent plate mounted on the polarimeter.)

If it continues to rotate excessively, follow previous procedure again.

Mark the spot where the laser hits the tape with a felt tip pen.

Take the mount off the polarimeter (that includes the taped exit hole)

Put the detector on the mount. Look at the detector entrance through the

mount to make sure that the mark is in the center of the detector opening.

When the detector is in place screw in the detector with the set screws so that

it stays in that spot.

Take the detector of with its support. Remove the tape and put the detector with its support back onto the mount. Make sure not to bump the set screws.