# Tecnai F30

The following procedures are meant to assist the new user to start a “normal use” microscope session. If there are questions with any of the steps, or if it appears there may be a problem with the microscope DO NOT TRY TO SOLVE IT YOURSELF—GET HELP.

## CHECKLIST Starting microscope operation

The checklist is meant as a reminder. If you need more than a reminder, read the sections referred to by the corresponding steps. FEEL FREE TO ASK FOR HELP! The Lab manager is here to assist and train users as well as keep the instrument in working condition, so he is willing to help.

- Check the lighted buttons on the right side of the column. The lights should be lit from left to right: RED-WHITE-WHITE. If it is other wise, see the Lab manager.

- Fill the cold trap with liquid nitrogen (For more info, see the section “Cold Trap.”)

- Log in to Windows using the username and password that was set up by the Lab manager during training. (For more info, see the section “Logging in.”)

- If necessary, start the “Tecnai user interface.” It should start up automatically.

- Check the Vacuum control panel. The column vacuum should be around 6 to 12 (log units) and the gun vacuum around 1. (For more info, see the section “Vacuum.”)

- Remove the specimen holder that is currently in the microscope by following the steps below. There will always be a holder in the microscope. (For more info, see the section “Removing a specimen holder from the microscope.”)

  - Ensure that the column valves are closed. In the Vacuum control panel the “Col Valves Closed” button should be YELLOW.

  - Go to the Stage control panel flap-out menu and press the “Reset the holder” button under the “control” tab. Wait for the tilts to go to zero.

  - Stand up.

  - Pull the specimen rod straight out until it stops (continue holding the rod or it will go back into the column).

  - Turn the specimen rod clockwise (do not continue pulling-JUST TURN) until it stops. Before performing the next step make sure the rod stays turned all the way clockwise.

  - Place your left hand on the base of the specimen exchange area and your right hand on the rod. Place your right thumb on your left hand and use your thumb to break the vacuum of the specimen exchange chamber.

  - After breaking the vacuum gently and steadily pull the rod completely out of the specimen exchange.

  - Remove the beta-tilt cable.

- Insert a specimen into a specimen holder (see instructions for this operation which should always be located near the holder).

- Insert the holder into the microscope by following the steps below (For more info, see the section “Inserting a specimen holder into the microscope.”)

  - Stand up

  - Before insertion, align the small pin near the end of the specimen rod with the slit on the front of the specimen exchange. The better they are aligned the easier this will be.

  - Slide the specimen rod into the specimen exchange until it comes to a stop. If you are at the right stop, the rod will not twist on its own. It will be “snug” to the touch. (It will be exactly 2 and 3/8 inches from the front of the specimen exchange.) The little red LED on the front of the specimen exchange will light up at this point.

  - CAREFULLY plug in the beta-tilt cable.

  - Select the specimen holder in the Tecnai User Interface (bottom middle) and select return.

  - Select return again for the dialogue “Connect B tilt cable.”

  - NOW WAIT FOR THE RED LED TO GO OFF BEFORE PROCEEDING.

  - Once the red LED has gone off turn the rod counter-clockwise SLOWLY 120 degrees. After 120 degree turn the rod will feel that it is being pulled into the column

  - Allow the rod to go into the column all the way. (If it is all the way in, you should not be able to fit your little finger behind it—unless you are a two-year old.)

  - Shut off the turbo pump by deselecting the “Turbo on” button on the Vacuum control panel.

- Wait for the column vacuum to drop below 12 (log units), then open the column valves. The beam should now be visible. (If the beam is not visible see the section “Finding the beam.”)

## CHECKLIST  Ending a microscope session

When you are finished working with the microscope follow the steps below, and remember that you will simply log off (in the Start menu, select Log Off). Do not select “shut down the computer” because that will also effectively shut down the microscope. If you only log off, the microscope will keep running normally (except that nobody can change settings until logged in).

Before logging off (or closing the user interface):

- Turn the magnification to a reasonable level (around 8000x).

- Expand the beam to fill the viewing screen.

- Close the column valves.

- Remove the specimen holder. (For more info, see the section “Removing a specimen holder from the microscope.”)

  - Ensure that the column valves are closed. In the Vacuum control panel the “Col Valves Closed” button should be YELLOW.

  - Go to the Stage control panel flap-out menu and press the “Reset the holder” button under the “control” tab. Wait for the tilts to go to zero.

  - Stand up.

  - Pull the specimen rod straight out until it stops (continue holding the rod or it will go back into the column).

  - Turn the specimen rod clockwise (do not continue pulling-JUST TURN) until it stops. Before performing the next step make sure the rod stays turned all the way clockwise.

  - Place your left hand on the base of the specimen exchange area and your right hand on the rod. Place your right thumb on your left hand and use your thumb to break the vacuum of the specimen exchange chamber.

  - After breaking the vacuum gently and steadily pull the rod completely out of the specimen exchange.

  - Remove the beta-tilt cable.

- Remove the specimen from the specimen holder (see instructions for this operation which should always be located near the holder).

- If there is nobody waiting to use the microscope immediately after you, please insert an empty specimen holder into the microscope by following the steps below (For more info, see the section “Inserting a specimen holder into the microscope.”)

  - Stand up

  - Before insertion, align the small pin near the end of the specimen rod with the slit on the front of the specimen exchange. The better they are aligned the easier this will be.

  - Slide the specimen rod into the specimen exchange until it comes to a stop. If you are at the right stop, the rod will not twist on its own. It will be “snug” to the touch. (It will be exactly 2 and 3/8 inches from the front of the specimen exchange.) The little red LED on the front of the specimen exchange will light up at this point.

  - CAREFULLY plug in the beta-tilt cable.

  - Select the specimen holder in the Tecnai User Interface (bottom middle) and select return.

  - Select return again for the dialogue “Connect B tilt cable.”

  - NOW WAIT FOR THE RED LED TO GO OFF BEFORE PROCEEDING.

  - Once the red LED has gone off turn the rod counter-clockwise SLOWLY 120 degrees. After 120 degree turn the rod will feel that it is being pulled into the column

  - Allow the rod to go into the column all the way. (If it is all the way in, you should not be able to fit your little finger behind it—unless you are a two-year old.)

  - Shut off the turbo pump by deselecting the “Turbo on” button on the Vacuum control panel.

- Exit the Tecnai user interface, and save changes when prompted.

- Log off the Windows interface.